OPTRIS MSpro/ LS

Communication interface

(Revision 1.5 / 01.09.2006)

1 Serial interface parameters

Baud rate:	115200

Data bits:	8

Parity:	none

Stop bits:	1

Flow control:	off

2 Protocol

The protocol of the OPTRIS MSpro/LS is a binary protocol. A checksum is needed for all commands. The protocol has no additional overhead with CR, LR or ACK bytes. This makes the communication fast. Any command starts with a single byte command code.

Read commands:	01hex .. 2Fhex

Set commands: 	81hex .. AFhex

Example: To get the current object temperature the user must send a simple 01hex byte (and checksum 01hex) and the device will respond with the two byte temperature. Deducted by 1000dec  the result is the temperature in tenth degrees. For more examples see last section of this document.

Set command byte must be followed by the appropriate data. For example to set the emissivity to 0.95 the user must send the set command byte A0hex , the data bytes 03hex B6hex , and the checksum 15hex.

3 Data formats

All numbers with more than one byte are big endian encoded (most significant byte first).

| Data | Bytes | Encoding | Example | Remarks |
| --- | --- | --- | --- | --- |
| Temperatures / VCC | 2 | (byte1*256 + byte2 - 1000) / 10 | 04 D3 -> = (04D3hex -1000)/10 = (1235-1000)/10 = 23.5°C | Any temperature on the interface is in Celsius. Conversion from/to Fahrenheit must be done by the PC. |
| Emissivity | 2 | (byte1*256 + byte2) / 1000 | 03 B6 -> = 03B6hex/1000 = 950/1000 = 0.950 |  |
| Energy Gain | 2 | (1/2^15) * (byte1*256 + byte2) | 81 93 -> (1/2^15) * 8193hex = 33171 / 32768 = 1.0123 |  |
| Firmware/ Hardware revision | 2 | Fixed point encoding: High byte = major Low byte = minor | 02 12-> = rev 2.18 |  |
| Serial number | 4 | unsigned long | 003DCC5D -> = 3DCC5Dhex = 4050013 |  |

4 Command list

The following table lists all commands accessible for the user.

| Function | Data bytes | R read | W write | Unit | Description / remarks | MS pro | LS |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Device |  |  |  |  | static device data |  |  |  |
| Firmware revision | 2 | 10 | - |  | See data format | x | x |  |
| Hardware revision | 2 | 11 | - |  | See data format | x | x |  |
| Serial no. | 4 | 12 | - |  | See data format | x | x |  |
| Production date | 3 | 13 | - |  |  | ? | ? |  |
| Calibration date | 3 | 14 | - |  |  | ? | ? |  |
| Lower limit of device range | 2 | 15 | - | °C | For example -40.0°C | x | x |  |
| Upper limit of device range | 2 | 16 | - | °C | For example 900.0°C | x | x |  |
| Model code | 1 | 17 | - |  | To distinguish different models | x | x |  |
| Code for device capabilities | 2 | 18 | - |  | for different capabilities | x | x |  |
| Setup |  |  |  |  | User device data |  |  |  |
| Emissivity | 2 | 20 | A0 | ε | Default = 0.95 | x | x |  |
| High alarm value | 2 | 21 | A1 | °C | Default = 100°C | x | x |  |
| Low alarm value | 2 | 22 | A2 | °C | Default = 0°C | x | x |  |
| Gain | 2 | 24 | A4 |  | External adjustable gain (energy) Default = 1.0 | x | x |  |
| Offset | 2 | 25 | A5 | °C | Ext. adjustable offset (temperature) Default = 0.0 | x | x |  |
| Ambient temperature | 2 | 26 | A6 | °C | Ambient-Control, Default = 23°C |  | x |  |
| Setup word | 2 | 27 | A7 |  | See table | x | x |  |
| Bargraph cycle time | 1 | 28 | A8 | mec. | Cycle time of the LS display’s bargraph Default = 40ms |  | x |  |
| Date / time | 4 | 29 | A9 | sec. | Seconds since 1.1.2001 00:00 | x | x |  |
| Data |  |  |  |  | Read only data |  |  |  |
| TObj | 2 | 01 | - | °C | Object temperature | x | x |  |
| TInt | 2 | 02 | - | °C | Internal temperature / head Temperature | x | x |  |
| TObj Max | 2 | 03 | - | °C | Maximum object temperature | x | x |  |
| TObj Min | 2 | 04 | - | °C | Minimum object temperature | x | x |  |
| TObj Avg | 2 | 05 | - | °C | Average object temperature | x | x |  |
| TTC | 2 | 06 | - | °C | TC probe temperature | x | x |  |
| TTC-Socket | 2 | 07 | - | °C | Temperature of the TC socket | x | x |  |
| Vcc | 2 | 0E | - | V | Battery voltage (debug only) |  | x |  |
| Control |  |  |  |  |  |  |  |  |
| Control byte | 1 | - | AD |  | Delete logger, set default: see control byte | x | x |  |
| Start boot loader | x |  | 3F |  | Firmware update only | x | x |  |
| Memory |  |  |  |  | Data logger |  |  |  |
| Used data logger entries | 2 | 19 | - |  | Number of used logger entries | x | x |  |
| Data logger entry | x | 1A | - |  | The command must follow the number of the logger entry 1A 00: live data burst (current temp’s) 1A 01: first logger entry | x | x |  |
| Misc. |  |  |  |  |  |  |  |  |
| LCD memory | 20 | 1C | - |  | For debug only |  | x |  |
| ADC-Werte | 16 | 1D | - |  | For debug only |  | x |  |
| General status | 2 | 1E | - |  | See status general word | x | x |  |
| Key status | 2 | 1F | 9F |  | See key status word. |  | x |  |
| Material table | x * 6 | 2B | AB |  | 2 bytes for emissivity / 4 bytes for material |  | x |  |
| Table for predefined data of logger remarks | x * 4 | 2C | AC |  | 4 bytes for any predefined remark |  | x |  |

5 Data structures

Setup word (R:27 / W:A7)

| Data | Bit | Default |  | MSpro | LS |  |
| --- | --- | --- | --- | --- | --- | --- |
| High Alarm | 0 | On | 1 = trigger high Alarm |  | x |  |
| Low Alarm | 1 | On | 1 = trigger low alarm |  | x |  |
| Temperature unit | 2 | °C | 0 = °C, 1 = °F | x | x |  |
| Audible alarm | 3 | On | 1 = Audible alarm available | x | x |  |
| TC | 4 | Off | 1 = TC-Temp. Available in display |  | x |  |
| Lock | 5 | Off | 1 = device do not turn off automatically | x | x |  |
| Disable user interface (key’s) | 6 | Off | 1 = user interface is locked |  | x |  |
| Ambient Source | 7 | TInt | 1 = TInt , 0 = user value (cmd. 0xA6) |  | x |  |
| LASER-Flash | 8 | Off | 1 = LASER flashes on alarm condition |  | x |  |
| Bargraph mode | 9 | 1 | 1 = Difference between Max und Min 0 = Difference between Hi/Lo-Alarm |  | x |  |
| Makro mode | 10,11 | Auto | 0 = Off, 1 = On, 2= Auto |  | x |  |
| Material table available | 12 | Off | 1 = besides emissivity the material names are usable |  | x |  |
| User definable remarks for data logger | 13 | Off | 1 = User definable remarks for data logger are available |  | x |  |
| LASER | 14 | On | 1 = LASER available | x | x |  |
| Backlight | 15 | On | 1 = Backlight available | x | x |  |

Control byte (W: AD)

| Command | Data |  | MSpro | LS |  |
| --- | --- | --- | --- | --- | --- |
| Delete data logger | 0x03 | Deletes all data logger entries | x | x |  |
| Set default | 0x05 | Set device to factory default | x | x |  |

General status word  (R:1E)

| Command | Bit |  | MSpro | LS |  |
| --- | --- | --- | --- | --- | --- |
| High alarm | 0 | True, if alarm is active | x | x |  |
| Low alarm | 1 | True, if low alarm is active | x | x |  |
| LASER | 3 | True, if LASER is on | x | x |  |
| Back light | 4 | True, if back light is on | x | x |  |
| CF/SF | 5 | 1 = Close focus, 0 = Standard focus |  | x |  |
| Tilt sensor | 6 | 1 = device is tilt |  | x |  |
| Logger full | 8 | All Logger entries are used | x | x |  |
| Thermocouple | 15 | True, if TC present | x | x |  |

Key status word (R:1F / W:9F)

Set command can be used to emulate a pressed key.

Note: Emulating the Trigger will not activate the LASER (for security)

| MSpro |  |  |
| --- | --- | --- |
| Command | Bit |  |
| Trigger | 0 | 1, if trigger is pushed |
| Up | 1 | 1, if Up is pushed |
| Down | 2 | 1, if Down is pushed |
| Mode | 3 | 1, if Mode is pushed |

| LS |  |  |
| --- | --- | --- |
| Command | Bit |  |
| Trigger | 0 | 1, if trigger is pulled |
| Up | 1 | 1, if Up is pushed |
| Down | 2 | 1, if Down is pushed |
| Mode1 | 3 | 1, if Mode1 is pushed |
| Mode2 | 4 | 1, if Mode2 is pushed |

6 Data logger

Byte sequences for command 1A:

| MS pro |  |  |
| --- | --- | --- |
| Value | Bytes | Unit |
| Entry number | 1 | - |
| TObj | 2 | °C |
| TInt | 2 | °C |
| TTC | 2 | °C |
|  | Σ=7 |  |

| LS |  |  |
| --- | --- | --- |
| Value | Bytes | Unit |
| Entry number | 1 | - |
| TObj | 2 | °C |
| TInt | 2 | °C |
| TObj Max | 2 | °C |
| TObj Min | 2 | °C |
| TObj Avg | 2 | °C |
| TTC | 2 | °C |
| High alarm | 2 | °C |
| Low alarm | 2 | °C |
| Emissivity | 2 | - |
| Time | 4 | sec. |
| Remark | 4 | ASCII |
|  | Σ=27 |  |

7 Checksums

Any command must have a checksum suffix. Any answer has a checksum suffix.

The checksum byte is build by the arithmetical XOR of all command bytes.

8 Examples

| Read Commands | Send *) | Receive*) |  |
| --- | --- | --- | --- |
| Reading a target temperature | 01 [01] | 04 D3 [D7] | (04D3hex -1000)/10 = (1235-1000)/10 = 23.5°C |
| Reading emissivity | 20 [20] | 03 B6 [B5] | 03B6hex/1000 = 950/1000 = 0.950 |
| Reading serial number | 12 [12] | 3D CC 5D | 3DCC5Dhex = 4050013 |
| Reading first data logger entry | 1A 01 [1B] | 04 D3 … [xx] | 22 bytes (MS/LS) |
| Reading status | 1E [1E] | 31 [31] | High alarm condition, LASER and back light is active |
| Reading key status | 1F [1F] | 08 [08] | LS: Mode1 is pushed |

| Set Commands | Send *) | Receive*) |  |
| --- | --- | --- | --- |
| Setting the high alarm value | A1 04 D3 [76] | 04 D3 [D7] | (04D3hex -1000)/10 = (1235-1000)/10 = 23.5°C |
| Setting the emissivity to 0.95 | A0 03 B6 [15] | 03 B6 [B5] | 03B6hex/1000 = 950/1000 = 0.950 |
| Delete data logger | AD 03 [AE] | 03 [03] | Deletes all data logger entries |
| Set default | AD 05 [A8] | 05 [05] | Set device to factory default |
| Push a key via ser. interface | 9F 04 [9B] | 04 [04] | Simulates push “Down” |

*) Checksum in square brackets